How to Use a VeraTemp thermometer

Each morning:

- Allow 15 minutes for the VeraTemp thermometer to reach the temperature of the area in which it is to be used.

- Look at the lens at the end of the VeraTemp. If it is dirty, clean it: dip a cotton swab in 70° isopropyl or rubbing alcohol and rub it gently over the lens. Do not use the VeraTemp until the lens is completely dry.

- Make sure the VeraTemp Mode is set to “Body.”  If you need to, press the mode button to switch between “Surface Temp,” “Body,” and “Room.”

Using the thermometer:

- Ask the traveler to remove any hat or glasses, push back his or her hair, and wipe off any sweat.

- Aim for the forehead, between the eyebrows and the hair.

- Hold the VeraTemp 5-8cm  (thickness of 3 fingers) from the person.  The reading will be wrong if the VeraTemp is held farther away. Do not touch the skin with the VeraTemp.

- Press the scan button to take the temperature. The temperature will be shown on the screen. The screen will be red when the temperature is higher than 38° C.

- If the reading is below 35.8° retake the temperature with the same VeraTemp.

- Before you retake the temperature, make sure there is nothing in the way (such as hair, glasses, hat), the target area is sweat-free, and the VeraTemp is 5-8 cm away from the forehead.

- If the second reading is below 35.8°C, retake the temperature with a different VeraTemp.

- If the temperature on the second VeraTemp is still below 35.8°C, record the reading if required on the correct form and note that you took the temperature more than once and with different devices.

- If the temperature is now 35.8°C or greater, record the new temperature if required on the right form.  Stop using the first VeraTemp until it can be checked.

- If the reading is 37.5°C or higher (or the threshold of the country), the person has a fever.

- Move the person at least 1 meter from other people and keep him or her there.

- If required, record the temperature on the right form.

- Alert your supervisor and follow procedures to ask for a medical evaluation.

Never touch the glass lens of the VeraTemp directly with your fingers. It is the most fragile part of the thermometer and may be damaged if you do.